## Foglio1
Mapping of new ERF8 board & Preamp. Calibration
Chain: Pulser->capacity 1pF->64ch ERF8 Adapter->Samtec cable->PA->digitizer
Calibration performed with CoMPASS. See at the right for pulser and digitizer settings | [14 empty] | Pulser settings |  |  | Digitizer settings
the first S of board input is the ground | [14 empty] | fall time 20us |  |  | as written in the first page of quaderno rosso dal 22/06/2022
PA 1 (21119) + DIG 1 (22642) | [5 empty] | PA1 with DIG2 (22643)+flat cables of "2" | [8 empty] | rate 100Hz
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated |  | E centroid (ch) | FWHM (ch) | difference |  |  | 14410.3125 | average |  |  | delay 150ns
S1 | 0 | 14338.5 | 14 | 1 |  | 14363 | 14 | 0.1709% |  |  | -4.096% | max variation from average |  |  | amp 2V
S2 | 16 | 14665.5 | 13 | 0.9777 |  | 14696 | 15 | 0.208% |  |  | 7.946% | variation max - min |  |  | polarity Negative
S3 | 1 | 14163 | 11 | 1.0124 |  | 14117 | 16 | -0.3248% | [6 empty] | pulse top tail
S4 | 17 | 14632 | 14 | 0.97994 |  | 14577 | 16 | -0.3759% | [6 empty] | atten 1X
S5 | 2 | 14261 | 11 | 1.0054 |  | 14253 | 16 | -0.0561%
S6 | 18 | 14728.5 | 12 | 0.97352 |  | 14755 | 17 | 0.1799%
S7 | 3 | 14658 | 13 | 0.9782 |  | 14609 | 18 | -0.3343%
S8 | 19 | 14709 | 14 | 0.97481 |  | 14724 | 19 | 0.102%
S9 | 4 | 14413 | 11 | 0.99483 |  | 14357 | 14 | -0.3885% | [8 empty] | What is the difference if we use "test input"-PA-DIG?
S10 | 20 | 14371.5 | 12 | 0.9977 |  | 14373 | 14 | 0.01044% | [8 empty] | ERF8 channel | Digitizer channel | E centroid (ch) | with respect to the previous results
S11 | 5 | 14227 | 16 | 1.0078 |  | 14208 | 13 | -0.1335% | [8 empty] | S1 | 0 | 16660 | 1.1619
S12 | 21 | 14700 | 12 | 0.97541 |  | 14634 | 13 | -0.449% | [8 empty] | S2 | 16 | 17179 | 1.1714
S13 | 6 | 14075 | 12 | 1.0187 |  | 14138 | 14 | 0.4476% | [8 empty] | S3 | 1 | 16326 | 1.1527
S14 | 22 | 14366 | 13 | 0.99809 |  | 14312 | 13 | -0.3759% | [8 empty] | S4 | 17 | 17757 | 1.2136
S15 | 7 | 14329.5 | 13 | 1.0006 |  | 14346 | 14 | 0.1151% | [8 empty] | S5 | 2 | 16549 | 1.1604
S16 | 23 | 14665 | 14 | 0.97774 |  | 14653.5 | 16 | -0.07842%
S17 | 8 | 14351 | 14 | 0.99913 |  | 14247 | 15 | -0.7247%
S18 | 24 | 14504 | 13 | 0.98859 |  | 14608 | 15 | 0.717% | [8 empty] | V=1.724 V (instead of 2V)=2/1.16 (with test input)
S19 | 9 | 14528.5 | 13 | 0.98692 |  | 14445 | 14 | -0.5747% | [8 empty] | ERF8 channel | Digitizer channel | E centroid (ch)
S20 | 25 | 14495 | 13 | 0.9892 |  | 14562 | 14 | 0.4622% | [8 empty] | S1 | 0 | 14335 | 0.99976
S21 | 10 | 14383 | 14 | 0.99691 |  | 14359 | 14 | -0.1669% | [8 empty] | S2 | 16 | 14779 | 1.0077
S22 | 26 | 14793 | 14 | 0.96928 |  | 14856 | 14 | 0.4259% | [8 empty] | S3 | 1 | 14047 | 0.99181
S23 | 11 | 14222 | 11 | 1.0082 |  | 14227 | 14 | 0.03516% | [8 empty] | S4 | 17 | 15278 | 1.0441
S24 | 27 | 14724.5 | 12 | 0.97379 |  | 14720 | 14 | -0.03056% | [8 empty] | S5 | 2 | 14237 | 0.99832
S25 | 12 | 14251 | 12 | 1.0061 |  | 14267 | 14 | 0.1123%
S26 | 28 | 14599 | 12 | 0.98216 |  | 14646 | 16 | 0.3219%
S27 | 13 | 14514 | 12 | 0.98791 |  | 14506 | 14 | -0.05512%
S28 | 29 | 14480 | 12 | 0.99023 |  | 14459 | 14 | -0.145% | [8 empty] | Is the calibration linear with intercept=0 ?
S29 | 14 | 14359 | 12 | 0.99857 |  | 14396 | 13 | 0.2577% | [8 empty] | for S1=ch0 of dig 1 | pulser precision:+/-10mV | NOT TEST INPUT
S30 | 30 | 14414 | 12 | 0.99476 |  | 14422.5 | 15 | 0.05897% | [8 empty] | Volt | centroid (test input) | centroid (ch) | large V value with attenuation to decrease error
S31 | 15 | 14419 | 14 | 0.99442 |  | 14378 | 16 | -0.2843% | [8 empty] | 1 | 8335 | 118 | 0.016
S32 | 31 | 14641 | 13 | 0.97934 |  | 14645 | 16 | 0.02732% | [8 empty] | 1.5 | 12500 | 293 | 0.04
S33 | 32 | 14260 | 11 | 1.0055 | [12 empty] | 2 | 16665 | 585 | 0.08
S34 | 48 | 14761 | 12 | 0.97138 | [12 empty] | 2.5 | 20835 | 1151 | 0.16
S35 | 33 | 14965 | 11 | 0.95814 | [12 empty] | 3 | 24996 | 2895 | 0.4
S36 | 49 | 14688 | 11 | 0.97621 | [12 empty] | 4 | 33306 | 5782 | 0.8
S37 | 34 | 14296 | 12 | 1.003 | [11 empty] | fit in root |  |  | 11356 | 1.6
S38 | 50 | 14448 | 11 | 0.99242 | [14 empty] | 28614 | 4
S39 | 35 | 14172 | 11 | 1.0117 | [14 empty] | 57018 | 8
S40 | 51 | 13959 | 12 | 1.0272 | [12 empty] | slope | 8324.83 +/- 3 | slope | 7128 +/- 7
S41 | 36 | 14307 | 12 | 1.0022 | [12 empty] | intercept | 14.9 +/- 7.4 | intercept | 23 +/- 20
S42 | 52 | 14440 | 11 | 0.99297
S43 | 37 | 14226 | 12 | 1.0079
S44 | 53 | 14417.5 | 12 | 0.99452
S45 | 38 | 14471 | 12 | 0.99084
S46 | 54 | 14607 | 13 | 0.98162
S47 | 39 | 14592 | 11 | 0.98263
S48 | 55 | 14263 | 13 | 1.0053
S49 | 40 | 14283 | 12 | 1.0039
S50 | 56 | 13820 | 12 | 1.0375
S51 | 41 | 14197 | 14 | 1.01
S52 | 57 | 14775 | 13 | 0.97046
S53 | 42 | 14302 | 11 | 1.0026
S54 | 58 | 14026 | 13 | 1.0223
S55 | 43 | 14160 | 12 | 1.0126
S56 | 59 | 14603 | 13 | 0.98189
S57 | 44 | 14415 | 12 | 0.99469
S58 | 60 | 14167 | 13 | 1.0121
S59 | 45 | 14510 | 12 | 0.98818
S60 | 61 | 14541 | 13 | 0.98607
S61 | 46 | 14158 | 12 | 1.0127
S62 | 62 | 13956 | 13 | 1.0274
S63 | 47 | 14086 | 13 | 1.0179
S64 | 63 | 14408 | 12 | 0.99518
*with some minutes/hours of difference without close anything or closing & starting pulser: ±(0-1) ch
*with better statistics (7k->70k): ±0 ch
*test for S2,S16,S17,S63 after a day and after restarting crates: +(3-8) ch and FWHM: ±(0-1) ch. After an hour the values approach more the ones of the previous day (perhaps a matter of temperature).
* S33-S64 were evaluated with a day of difference (with respect to S1-S32)
* after the afformentioned test with a different digitizer, it was tested that the original centroids with the original chain were invariant: E ±(0-2) ch
PA 2 (21114)+ DIG 2 (22643)
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated |  |  |  | 14263.0156 | average
S1 | 0 | 14392 | 15 | 0.99628 |  |  |  | -3.653% | max variation from average
S2 | 16 | 14053 | 14 | 1.0203 |  |  |  | 7.172% | variation max - min
S3 | 1 | 14196 | 15 | 1.01
S4 | 17 | 14020 | 14 | 1.0227
S5 | 2 | 14222 | 13 | 1.0082
S6 | 18 | 14275 | 14 | 1.0044
S7 | 3 | 14193 | 13 | 1.0103
S8 | 19 | 13939 | 16 | 1.0287
S9 | 4 | 14296 | 15 | 1.003
S10 | 20 | 13903 | 12 | 1.0313
S11 | 5 | 14584 | 13 | 0.98317
S12 | 21 | 14171 | 13 | 1.0118
S13 | 6 | 14565 | 12 | 0.98445
S14 | 22 | 13925 | 14 | 1.0297
S15 | 7 | 14536 | 14 | 0.98641
S16 | 23 | 14091.5 | 13 | 1.0175
S17 | 8 | 14314 | 14 | 1.0017
S18 | 24 | 13972 | 12 | 1.0262
S19 | 9 | 14320 | 11 | 1.0013
S20 | 25 | 14358 | 12 | 0.99864
S21 | 10 | 14211 | 13 | 1.009
S22 | 26 | 14385 | 13 | 0.99677
S23 | 11 | 14506 | 12 | 0.98845
S24 | 27 | 14346 | 15 | 0.99948
S25 | 12 | 14765 | 12 | 0.97111
S26 | 28 | 13974.5 | 15 | 1.026
S27 | 13 | 14309 | 16 | 1.0021
S28 | 29 | 14344 | 13 | 0.99962
S29 | 14 | 14234 | 11 | 1.0073
S30 | 30 | 14157 | 14 | 1.0128
S31 | 15 | 14261 | 10 | 1.0054
S32 | 31 | 14713 | 12 | 0.97455
S33 | 32 | 14413 | 12 | 0.99483
S34 | 48 | 14021 | 12 | 1.0226
S35 | 33 | 14608 | 12 | 0.98155
S36 | 49 | 14184 | 12 | 1.0109
S37 | 34 | 14619.5 | 11 | 0.98078
S38 | 50 | 14394 | 12 | 0.99614
S39 | 35 | 14741 | 13 | 0.9727
S40 | 51 | 14002 | 11 | 1.024
S41 | 36 | 14607 | 12 | 0.98162
S42 | 52 | 14392 | 12 | 0.99628
S43 | 37 | 14488 | 12 | 0.98968
S44 | 53 | 14285 | 11 | 1.0037
S45 | 38 | 14265 | 11 | 1.0052
S46 | 54 | 14226 | 12 | 1.0079
S47 | 39 | 14498 | 12 | 0.989
S48 | 55 | 14218 | 12 | 1.0085
S49 | 40 | 14248 | 14 | 1.0064
S50 | 56 | 14022 | 12 | 1.0226
S51 | 41 | 14523 | 12 | 0.9873
S52 | 57 | 14366 | 13 | 0.99809
S53 | 42 | 14175 | 12 | 1.0115
S54 | 58 | 14027 | 11 | 1.0222
S55 | 43 | 14114.5 | 12 | 1.0159
S56 | 59 | 14101 | 13 | 1.0168
S57 | 44 | 14262 | 14 | 1.0054
S58 | 60 | 14083 | 12 | 1.0181
S59 | 45 | 14278 | 14 | 1.0042
S60 | 61 | 13742 | 14 | 1.0434
S61 | 46 | 14049 | 13 | 1.0206
S62 | 62 | 13941 | 13 | 1.0285
S63 | 47 | 14346 | 13 | 0.99948
S64 | 63 | 14063 | 14 | 1.0196
*test for S10-S13,S16 with some minutes of difference closing & starting again pulser: ±(0-2) ch BUT, FWHM: 17-19->12-13
*test for S10-S13,S16 after an hour closing & starting again pulser: ±0 ch and FWHM: ±0 ch
*test for S15, S27, S31 with some minutes of difference closing & starting again pulser: ±(0-1) ch and FWHM: ±(0-1) ch
*test for S63, S64 with an hour of difference closing & starting again pulser: ±3 ch and FWHM: ±0 ch
*test for S1,S16,S64 after a day and after restarting crates: +(3-8) ch and FWHM: ±(0-1) ch.
*test for S1,S16,S64 8 hours after the test mentioned in bold above: ±0 ch and FWHM: -(1-2) ch with respect to test above (PA power supply was switched off for 1 hour and now the PA surface temperature is OK).
* S33-S64 were evaluated with a day of difference (with respect to S1-S32) and in the meanwhile the power supply was removed and placed again (to decrease temperature)
PA 3 (21113)+ DIG 3 (22644) |  | VALUES IN RED were obtained in 28/07/2022 with the cable already in bad shape (see logbook) and therefore centroid with BAD PRECISION
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated | [5 empty] | TEST IN chain "2" |  |  | 14481.3415 | average
S1 | 0 |  |  | #DIV/0! | No signal, (but a small crosstalk to ch16,1 with small thresh) |  |  |  |  | NO SIGNAL |  |  | -3.434% | max variation from average
S2 | 16 | 14596 | 14 | 0.98236 | [5 empty] | OK |  |  | 5.98% | variation max - min
S3 | 1 | 14209 | 15 | 1.0091 | [5 empty] | OK
S4 | 17 | 14153 | 14 | 1.0131 | [5 empty] | OK
S5 | 2 |  |  | #DIV/0! | No signal |  |  |  |  | NO SIGNAL
S6 | 18 | 14515 | 14 | 0.98784 | [5 empty] | OK
S7 | 3 | 14616.5 | 14 | 0.98098 | [5 empty] | OK
S8 | 19 | 14397 | 13 | 0.99594 | [5 empty] | OK
S9 | 4 | 14398 | 14 | 0.99587 | [5 empty] | OK
S10 | 20 | 14094.5 | 13 | 1.0173 | [5 empty] | OK
S11 | 5 |  |  | #DIV/0! | No signal |  |  |  |  | NO SIGNAL
S12 | 21 | 13984.5 | 13 | 1.0253 | [5 empty] | OK
S13 | 6 | 14130 | 13 | 1.0148 | [5 empty] | OK
S14 | 22 | 14540 | 13 | 0.98614 | [5 empty] | OK
S15 | 7 | 14450.5 | 14 | 0.99225 | [5 empty] | OK
S16 | 23 | 14535.5 | 13 | 0.98645 | [5 empty] | OK
S17 | 8 | 14082 | 13 | 1.0182 | [5 empty] | OK
S18 | 24 | 14527 | 14 | 0.98702 | [5 empty] | OK
S19 | 9 | 14595 | 13 | 0.98243 | [5 empty] | OK
S20 | 25 | 14624 | 12 | 0.98048 | [5 empty] | OK
S21 | 10 | 14245 | 14 | 1.0066 | [5 empty] | OK
S22 | 26 | 14524 | 13 | 0.98723 | [5 empty] | OK
S23 | 11 | 14572 | 14 | 0.98398 | [5 empty] | OK
S24 | 27 | 14823 | 14 | 0.96731 | [5 empty] | OK
S25 | 12 | 14472 | 13 | 0.99078 | [5 empty] | OK
S26 | 28 | 14629 | 13 | 0.98014 | [5 empty] | OK
S27 | 13 | 14591 | 14 | 0.98269 | [5 empty] | OK
S28 | 29 | 14412 | 13 | 0.9949 | [5 empty] | OK
S29 | 14 |  |  | #DIV/0! | No signal |  |  |  |  | NO SIGNAL
S30 | 30 | 14843 | 13 | 0.96601 | [5 empty] | OK
S31 | 15 |  |  | #DIV/0! | No signal |  |  |  |  | NO SIGNAL
S32 | 31 | 14516 | 12 | 0.98777 | [5 empty] | OK
S33 | 32 | 14358 | 14 | 0.99864 | [5 empty] | OK
S34 | 48 | 14618 | 500 | 0.98088 | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S35 | 33 | 14603 | 20 | 0.98189 | [5 empty] | OK
S36 | 49 | 14337 | 500 | 1.0001 | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S37 | 34 | 15008 | 13 | 0.95539 | [5 empty] | OK
S38 | 50 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S39 | 35 |  |  | #DIV/0! | OK |  |  |  |  | OK
S40 | 51 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S41 | 36 |  |  | #DIV/0! | OK |  |  |  |  | OK
S42 | 52 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S43 | 37 |  |  | #DIV/0! | OK |  |  |  |  | OK
S44 | 53 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S45 | 38 |  |  | #DIV/0! | OK |  |  |  |  | OK
S46 | 54 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S47 | 39 | 14209 | 15 | 1.0091 | [5 empty] | OK
S48 | 55 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S49 | 40 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S50 | 56 | 14507 | 14 | 0.98838 | [5 empty] | OK
S51 | 41 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S52 | 57 | 14636 | 14 | 0.97967 | [5 empty] | OK
S53 | 42 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S54 | 58 | 14397 | 14 | 0.99594 | [5 empty] | OK
S55 | 43 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S56 | 59 | 14544.5 | 15 | 0.98584 | [5 empty] | OK
S57 | 44 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S58 | 60 | 14850 | 14 | 0.96556 | [5 empty] | OK
S59 | 45 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S60 | 61 | 14511 | 15 | 0.98811 | [5 empty] | OK
S61 | 46 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S62 | 62 | 14254 | 15 | 1.0059 | [5 empty] | OK
S63 | 47 |  |  | #DIV/0! | no signal at the beginning, OK on 28/07 |  |  |  |  | OK
S64 | 63 | 14828 | 15 | 0.96699 | [5 empty] | OK
* test at S2 with 12 hours of difference closing & starting again pulser: ±0 ch and FWHM: ±0 ch
* test at S2,S3,S8,S14 with a few minutes of difference and room temperature lower: for all of them centroid-> -5 ch and FWHM: ±(0-1) ch.
* test at S2,S28 with an hour of difference closing & starting again pulser and putting a cable for ch48-63: ±1 ch and FWHM: ±1 ch
* test for S2,S17,S52,S64 after restarting crates: ±(0-2) ch and FWHM: ±(0-1) ch.
* due to lack of cables, the same cables was used for 2 sets fo channels of pa3->dig3.
*VALUES IN RED were obtained in 28/07/2022 with the cable already in bad shape (see logbook) and therefore centroid with BAD PRECISION
PA 4 + DIG 4 (22645)
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated |  |  |  | 14379.2266 | average
S1 | 0 | 14237 | 13 | 1.0071 |  |  |  | -5.767% | max variation from average
S2 | 16 | 14285 | 13 | 1.0037 |  |  |  | 9.785% | variation max - min
S3 | 1 | 14408 | 12 | 0.99518
S4 | 17 | 14583 | 13 | 0.98323
S5 | 2 | 14332 | 13 | 1.0005
S6 | 18 | 14735 | 14 | 0.97309
S7 | 3 | 14043 | 13 | 1.021
S8 | 19 | 13898 | 13 | 1.0317
S9 | 4 | 14353.5 | 13 | 0.99895
S10 | 20 | 14728 | 14 | 0.97355
S11 | 5 | 14010 | 14 | 1.0234
S12 | 21 | 14601 | 14 | 0.98202
S13 | 6 | 14373 | 14 | 0.9976
S14 | 22 | 14460 | 13 | 0.9916
S15 | 7 | 14105 | 13 | 1.0166
S16 | 23 | 14502 | 13 | 0.98873
S17 | 8 | 14544 | 12 | 0.98587
S18 | 24 | 14567 | 13 | 0.98431
S19 | 9 | 14738 | 14 | 0.97289
S20 | 25 | 14552 | 14 | 0.98533
S21 | 10 | 14224 | 13 | 1.008
S22 | 26 | 14618 | 13 | 0.98088
S23 | 11 | 14323 | 14 | 1.0011
S24 | 27 | 14459 | 14 | 0.99167
S25 | 12 | 14207 | 14 | 1.0093
S26 | 28 | 14755 | 14 | 0.97177
S27 | 13 | 14370 | 13 | 0.99781
S28 | 29 | 14762.5 | 16 | 0.97128
S29 | 14 | 14262 | 15 | 1.0054
S30 | 30 | 14884 | 16 | 0.96335
S31 | 15 | 14418 | 14 | 0.99449
S32 | 31 | 14421 | 17 | 0.99428
S33 | 32 | 14366 | 12 | 0.99809
S34 | 48 | 14404 | 13 | 0.99545
S35 | 33 | 14219.5 | 12 | 1.0084
S36 | 49 | 14851 | 15 | 0.96549
S37 | 34 | 14477 | 14 | 0.99043
S38 | 50 | 13999 | 14 | 1.0243
S39 | 35 | 14197 | 13 | 1.01
S40 | 51 | 14624 | 14 | 0.98048
S41 | 36 | 14569 | 12 | 0.98418
S42 | 52 | 14395 | 15 | 0.99608
S43 | 37 | 14312 | 14 | 1.0019
S44 | 53 | 14957 | 14 | 0.95865
S45 | 38 | 14039.5 | 13 | 1.0213
S46 | 54 | 13550 | 14 | 1.0582
S47 | 39 | 14210.5 | 16 | 1.009
S48 | 55 | 14638 | 16 | 0.97954
S49 | 40 | 14098 | 16 | 1.0171
S50 | 56 | 14284 | 15 | 1.0038
S51 | 41 | 14356 | 14 | 0.99878
S52 | 57 | 13984 | 14 | 1.0254
S53 | 42 | 14646 | 18 | 0.979
S54 | 58 | 14401 | 17 | 0.99566
S55 | 43 | 14315 | 16 | 1.0016
S56 | 59 | 14448 | 17 | 0.99242
S57 | 44 | 14224 | 17 | 1.008
S58 | 60 | 14651 | 17 | 0.97867
S59 | 45 | 13891 | 16 | 1.0322
S60 | 61 | 14237 | 18 | 1.0071
S61 | 46 | 14235 | 16 | 1.0073
S62 | 62 | 14254 | 16 | 1.0059
S63 | 47 | 14183 | 15 | 1.011
S64 | 63 | 14497 | 18 | 0.98907
* test for S38,S44,S47,S62 with a few seconds of difference restarting or not restarting pulser: ±(0-1) ch and FWHM: ±(0-2) ch.
* test for S1 with 2 hours of difference: -4 ch and FWHM: +4 ch (PA surface has not a high temperature)
* test for S16 with 2 hours of difference: -6 ch and FWHM: +5 ch (PA surface has not a high temperature)
PA 5 + DIG 5 (22646)
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated |  |  |  | 14423.3125 | average
S1 | 0 | 14211.5 | 15 | 1.0089 |  |  |  | 3.499% | max variation from average
S2 | 16 | 14461 | 14 | 0.99153 |  |  |  | 6.157% | variation max - min
S3 | 1 | 14613 | 14 | 0.98122
S4 | 17 | 14474 | 14 | 0.99064
S5 | 2 | 14354 | 15 | 0.99892
S6 | 18 | 14676 | 15 | 0.977
S7 | 3 | 14367 | 14 | 0.99802
S8 | 19 | 14871 | 13 | 0.96419
S9 | 4 | 14485 | 13 | 0.98989
S10 | 20 | 14641 | 14 | 0.97934
S11 | 5 | 14071 | 13 | 1.019
S12 | 21 | 14678 | 15 | 0.97687
S13 | 6 | 14410 | 14 | 0.99504
S14 | 22 | 14494 | 14 | 0.98927
S15 | 7 | 14354 | 14 | 0.99892
S16 | 23 | 14169 | 14 | 1.012
S17 | 8 | 14299 | 18 | 1.0028
S18 | 24 | 14181 | 14 | 1.0111
S19 | 9 | 14596 | 14 | 0.98236
S20 | 25 | 14928 | 15 | 0.96051
S21 | 10 | 14413 | 12 | 0.99483 | at the beginning peak destroyed, rate 1.5kHz or a few mHz! (pulser 100 Hz). 1hour later ok !!! but with kHz in ch22,23
S22 | 26 | 14484 | 15 | 0.98995
S23 | 11 | 14412 | 12 | 0.9949
S24 | 27 | 14174 | 13 | 1.0116 | rate 300 Hz but peak ok (pulser 100 Hz). 1 hour later peak the same, rate ok, but 150kHz in ch 22
S25 | 12 | 14229 | 12 | 1.0077 | at the beginning peak destroyed, rate 1.5kHz or a few mHz! (pulser 100 Hz). 1hour later ok !!! but with kHz in ch22,23
S26 | 28 | 14450 | 14 | 0.99228
S27 | 13 | 14217 | 14 | 1.0085 | 96Hz but,for seconds rate 120 Hz ! Peak ok (pulser 100 Hz). Also, at the same moment 200Hz at ch 12
S28 | 29 | 14378 | 12 | 0.99725 | at the same moment (high) rate at ch 12, 27
S29 | 14 | 14373 | 13 | 0.9976 | at the same moment (high) rate at ch 11, 27
S30 | 30 | 14040 | 13 | 1.0213 | at the same moment a low rate also at ch 25
S31 | 15 | 14518 | 13 | 0.98764 | at the same moment an 120Hz rate at ch 25
S32 | 31 | 14672 | 13 | 0.97727 | at the same moment a 2 MHz rate at ch 9 ! The next day was ok without rate at ch 9
S33 | 32 | 14714 | 13 | 0.97448
S34 | 48 | 14528 | 14 | 0.98696
S35 | 33 | 14410 | 12 | 0.99504
S36 | 49 | 14739 | 13 | 0.97283
S37 | 34 | 14364 | 14 | 0.99822
S38 | 50 | 14102 | 12 | 1.0168 | at the same moment an 230Hz rate at ch 25 with E=0
S39 | 35 | 14663 | 13 | 0.97787 | at the same moment an 25kHz rate at ch 25 with E=0 (it needs a thr=6000 to be cutted)
S40 | 51 | 14168 | 13 | 1.012
S41 | 36 | 14643 | 11 | 0.97921
S42 | 52 | 14289 | 13 | 1.0035
S43 | 37 | 14242 | 12 | 1.0068
S44 | 53 | 14373 | 14 | 0.9976
S45 | 38 | 14456 | 14 | 0.99187
S46 | 54 | 14569 | 16 | 0.98418
S47 | 39 | 14595.5 | 14 | 0.98239
S48 | 55 | 14414 | 14 | 0.99476
S49 | 40 | 14311 | 12 | 1.0019
S50 | 56 | 14447 | 14 | 0.99249
S51 | 41 | 14372 | 14 | 0.99767
S52 | 57 | 14457 | 14 | 0.9918
S53 | 42 | 14326 | 14 | 1.0009
S54 | 58 | 14551 | 15 | 0.9854
S55 | 43 | 14248 | 15 | 1.0064
S56 | 59 | 14465 | 16 | 0.99125
S57 | 44 | 14751 | 16 | 0.97204
S58 | 60 | 14344 | 15 | 0.99962
S59 | 45 | 14272 | 14 | 1.0047
S60 | 61 | 14888 | 18 | 0.96309 | no signal the first day while a day after with destroyed signal, a few second after ok
S61 | 46 | 14098 | 13 | 1.0171
S62 | 62 | 14325 | 16 | 1.0009
S63 | 47 | 14060 | 16 | 1.0198
S64 | 63 | 14214 | 16 | 1.0088 | no signal the first day but a day after ok…
* test at S1 with 2 hours of difference closing & starting again pulser: ±0 ch and FWHM: ±1 ch. For some seconds random rate (mHz,Hz,kHz,MHz) at ch 9
* test at S19 (ch9) with 2 hours of difference closing & starting again pulser: -6 ch and FWHM: ±1 ch. No noise to this or to other channels.
* test with closed pulser: 800 Hz @ ch9 (cable connected to S32 / ch 31)
* test with closed pulser: 170 kHz @ ch24 (cable connected to S19 / ch 9). The rate appeared after several seconds.
* S33-S64: at ch with no relevant indication, no noise is appeared to any channel (it was checked with caution)
* test at S60 (ch61) 15 minutes later: again no signal but 4 kHz at ch9. A few sec later no rate at ch9.
* test at S64 (ch63) 5 minutes later: no signal to any channel
* test at S31 (ch15) an hour later: ±0 ch and FWHM: ±0 ch. No signal to any other channel
* S64 (ch63) was tested again after I saw again ch10,12 (in which kHz @ch22,23): again no signal, but kHz @ch 22,23 ! Same without pulser.
with the chain 3
PA (21116)+ DIG 3 (22644)
ERF8 channel | Digitizer channel | E centroid (ch) | FWHM (ch) | Calibrated |  | with chain "2" |  | 14461.3636 | average
S1 | 0 | 14235 |  | 1.0073 |  | OK |  | -3.301% | max variation from average
S2 | 16 | 14629 |  | 0.98014 |  | OK |  | 5.988% | variation max - min
S3 | 1 | 14383 | 32 | 0.99691 |  | OK
S4 | 17 | 14412 | 31 | 0.9949 |  | OK
S5 | 2 | 14253 |  | 1.006 |  | OK
S6 | 18 | 14685 |  | 0.9764 |  | OK
S7 | 3 | 14318 |  | 1.0014 |  | OK
S8 | 19 | 14561 |  | 0.98472 |  | OK
S9 | 4 | 14319 | 29 | 1.0014 |  | OK
S10 | 20 | 15004 |  | 0.95565 |  | OK
S11 | 5 | 14276 |  | 1.0044 |  | OK
S12 | 21 |  |  | #DIV/0! |  | OK
S13 | 6 |  |  | #DIV/0! |  | OK
S14 | 22 |  |  | #DIV/0! |  | OK
S15 | 7 |  |  | #DIV/0! |  | OK
S16 | 23 |  |  | #DIV/0! |  | OK
S17 | 8 |  |  | #DIV/0! |  | OK
S18 | 24 |  |  | #DIV/0! |  | OK
S19 | 9 |  |  | #DIV/0! |  | OK
S20 | 25 |  |  | #DIV/0! |  | OK
S21 | 10 |  |  | #DIV/0! |  | OK
S22 | 26 |  |  | #DIV/0! |  | OK
S23 | 11 |  |  | #DIV/0! |  | OK
S24 | 27 |  |  | #DIV/0! |  | OK
S25 | 12 |  |  | #DIV/0! |  | OK
S26 | 28 |  |  | #DIV/0! |  | OK
S27 | 13 |  |  | #DIV/0! |  | OK
S28 | 29 |  |  | #DIV/0! |  | OK
S29 | 14 |  |  | #DIV/0! |  | OK
S30 | 30 |  |  | #DIV/0! |  | OK
S31 | 15 |  |  | #DIV/0! |  | OK
S32 | 31 |  |  | #DIV/0! |  | OK
S33 | 32 |  |  | #DIV/0! |  | OK
S34 | 48 |  |  | #DIV/0! | no signal | OK
S35 | 33 |  |  | #DIV/0! |  | OK
S36 | 49 |  |  | #DIV/0! | no signal | OK
S37 | 34 |  |  | #DIV/0! |  | OK
S38 | 50 |  |  | #DIV/0! | no signal | OK
S39 | 35 |  |  | #DIV/0! |  | OK
S40 | 51 |  |  | #DIV/0! | no signal | OK
S41 | 36 |  |  | #DIV/0! |  | OK
S42 | 52 |  |  | #DIV/0! | no signal | OK
S43 | 37 |  |  | #DIV/0! |  | OK
S44 | 53 |  |  | #DIV/0! | no signal | OK
S45 | 38 |  |  | #DIV/0! |  | OK
S46 | 54 |  |  | #DIV/0! | no signal | OK
S47 | 39 |  |  | #DIV/0! |  | OK
S48 | 55 |  |  | #DIV/0! | no signal | OK
S49 | 40 |  |  | #DIV/0! | no signal | OK
S50 | 56 |  |  | #DIV/0! |  | OK
S51 | 41 |  |  | #DIV/0! | no signal | OK
S52 | 57 |  |  | #DIV/0! |  | OK
S53 | 42 |  |  | #DIV/0! | no signal | OK
S54 | 58 |  |  | #DIV/0! |  | OK
S55 | 43 |  |  | #DIV/0! | no signal | OK, but at the same time noise at ch 33,34,49
S56 | 59 |  |  | #DIV/0! |  | OK, but at the same time noise at ch 33,34,49
S57 | 44 |  |  | #DIV/0! | no signal | OK
S58 | 60 |  |  | #DIV/0! |  | OK
S59 | 45 |  |  | #DIV/0! | no signal | OK
S60 | 61 |  |  | #DIV/0! |  | OK
S61 | 46 |  |  | #DIV/0! | no signal | OK
S62 | 62 |  |  | #DIV/0! |  | OK
S63 | 47 |  |  | #DIV/0! | no signal | OK
S64 | 63 |  |  | #DIV/0! |  | OK
* from 27/07/2022 we saw at different chains that the peak is not very good (tail on the left side). The centroid is not significantly affected while the FWHM is higher. Perhaps the problem is the cable?